Laboratory Gap Analysis Guidance

Method for conducting an evidence-based ISO 15189 gap assessment

ISO 15189:2022 — Medical laboratories — Requirements for quality and competence

| Document Code | LQMS-GDE-002 | Version | 1.0 |
| --- | --- | --- | --- |
| Effective Date | 2026-09-19 | Classification | INTERNAL |
| Document Owner | [DOCUMENT OWNER] | Approver | [APPROVER] |
| Organization | [ORGANIZATION NAME] | Review Cycle | [ANNUAL / UPON CHANGE] |
| Website | www.iso-toolkits.org | Email | info@iso-toolkits.org |

[Note] Replace organization-specific placeholders and confirm applicable statutory, regulatory, accreditation and payer requirements before controlled release.

# 1. Purpose

This guidance provides a consistent method for assessing the current state of a medical laboratory against the intended requirements and operating outcomes of ISO 15189:2022. The objective is to identify evidence-based gaps, prioritize them according to risk and implementation impact, and convert them into a manageable action plan.

# 2. Assessment Principles

Assess actual practice and evidence, not only the existence of documents.

Use representative sampling across disciplines, shifts, sites, technologies and process stages.

Separate conformity gaps from improvement opportunities and implementation preferences.

Prioritize patient-safety, competence and result-reliability gaps over cosmetic documentation issues.

Record sufficient evidence so another competent reviewer can understand the basis of the conclusion.

# 3. Preparation

| Preparation activity | Practical output |
| --- | --- |
| Define scope | Sites, disciplines, POCT, support functions and exclusions to be assessed. |
| Identify assessors | Competent personnel with appropriate independence and technical support. |
| Collect documents | Policies, manual, procedures, method files, competence records, QC/EQA, audits, CAPA, review records. |
| Create sampling plan | Processes, records, staff interviews, equipment and examination files to sample. |
| Confirm logistics | Schedule interviews/observations without disrupting patient services. |

# 4. Evidence Collection Methods

Document review: confirm approved content, version, ownership and actual use.

Interview: test staff understanding of responsibilities, escalation and technical decisions.

Observation: compare documented processes with actual workflow and controls.

Record sampling: trace competence, maintenance, QC, EQA, nonconformity, critical results and other evidence.

Process tracing: follow a specimen or result end-to-end to test interface control and traceability.

# 5. Assessment Status Model

| Status | Use when | Typical action |
| --- | --- | --- |
| Fully Implemented | The process is defined, implemented, evidenced and functioning effectively. | Maintain and monitor. |
| Partially Implemented | Some controls exist but coverage, consistency, evidence or effectiveness is incomplete. | Complete or strengthen the control. |
| Not Implemented | The required control or reliable evidence is absent. | Create and implement a prioritized action. |
| N/A | The criterion is not applicable to the defined laboratory scope and the rationale is defensible. | Record the rationale. |
| Not Assessed | Evidence has not yet been reviewed. | Complete assessment before drawing a conclusion. |

# 6. Gap Analysis Workflow

Start with governance, scope and service profile so the assessment boundaries are clear.

Assess structural and resource controls, including competence, facilities, equipment, reagents, external services and information systems.

Assess pre-examination, examination and post-examination processes using end-to-end evidence tracing.

Assess management-system controls including document control, risk, nonconformity, complaints, continuity, indicators, audit and management review.

Record each gap in clear operational language: what is missing or ineffective, the evidence observed, and the consequence or risk.

Prioritize and assign actions; avoid combining unrelated causes into one vague action.

Validate conclusions with relevant process owners before finalizing the report.

Track actions to objective closure and reassess significant areas before accreditation readiness is declared.

# 7. Illustrative Gap Statements

[Note] The following examples are Illustrative Data only. They demonstrate writing quality, not findings about any laboratory.

| Example ID | Evidence observed | Assessment | Example gap statement | Priority logic |
| --- | --- | --- | --- | --- |
| EX-GAP-01 | Method verification exists for two analyzers; one recently introduced method has only the manufacturer insert in its file. | Partially Implemented | The laboratory has not documented verification that the new method performs acceptably under local conditions before routine reporting. | High: result reliability could be affected. |
| EX-GAP-02 | Competence assessments are recorded annually, but result authorization is not linked to documented authorization status. | Partially Implemented | The laboratory cannot consistently demonstrate that only currently authorized personnel release results. | High: governance and patient-safety impact. |
| EX-GAP-03 | Sample rejection reasons are recorded, but trends are not periodically analyzed. | Partially Implemented | Specimen-quality data are collected but are not used systematically to identify recurring pre-examination risks. | Medium: improvement opportunity with patient-service impact. |
| EX-GAP-04 | Critical results are phoned to clinical areas; acknowledgement time is not always captured. | Partially Implemented | Critical-result communication is not consistently traceable to recipient acknowledgement and time. | High: delay or miscommunication may affect care. |
| EX-GAP-05 | Internal audits cover document control and complaints but have not sampled technical examination processes. | Not Implemented | The audit programme does not yet provide assurance across the laboratory’s technical processes. | Medium/High depending on technical risk and maturity. |

# 8. Prioritization

Prioritization should combine the potential impact on patients or result reliability, regulatory/accreditation significance, breadth of exposure, recurrence, detectability and implementation dependency. A high-priority gap is one that warrants accelerated control, not simply one with extensive documentation work.

# 9. Reporting

Executive summary: scope, overall readiness themes, strengths and priority risks.

Detailed gap register: criterion, evidence, status, gap statement, priority and action owner.

Action plan: specific deliverables, accountable owner, due date, dependencies and closure evidence.

Limitations: areas not assessed, unavailable evidence and sampling constraints.

Reassessment plan: timing and criteria for verifying closure of significant gaps.

# 10. Common Failure Modes

Scoring without evidence or based only on staff statements.

Treating the presence of a procedure as proof the process is implemented.

Writing vague gaps such as “procedure inadequate” without describing the missing control.

Using one generic action such as “update SOP” where training, authorization, system configuration or technical verification is also required.

Declaring readiness while high-risk technical or patient-safety gaps remain open.

# References

| Reference | Title | Official source |
| --- | --- | --- |
| ISO 15189:2022 | Medical laboratories — Requirements for quality and competence | https://www.iso.org/standard/76677.html |
| ISO 20658:2023 | Requirements for the collection and transport of samples for medical laboratory examinations | https://www.iso.org/standard/80035.html |
| ISO 22367:2026 | Medical laboratories — Application of risk management to medical laboratories | https://www.iso.org/standard/87738.html |

[Note] Use the organization’s authorized copy of the applicable standard as the definitive reference. This document paraphrases requirements and does not reproduce the standard text.

# Revision and Approval

| Version | Date | Description | Approval Status |
| --- | --- | --- | --- |
| 1.0 | 2026-09-19 | Initial issue | Approved / [STATUS] |
| Prepared by | Reviewed by | Approved by | Approval Date |
| [DOCUMENT OWNER] | [REVIEWER] | [APPROVER] | [DATE] |